## Slide 1: Objective: To develop a standardised method for measuring annealing/ melting rates for complementary DNA oligomers by Fluorescence Spectroscopy.
Briefly,
  Complementary oligomers are mixed together in a standard buffer in equal molarity. A dsDNA specific dye, EvaGreen is added to the reaction mix to monitor the amount a dsDNA present at any given instant of time.
  When EvaGreen binds to dsDNA it is highly fluorescent. Thus, the fluorescent signal of a sample is directly proportional to and therefore a measure of the amount of dsDNA present in the sample.
  As the temperature of the sample is decreased, the complementary oligomers anneal resulting in the formation of dsDNA seen as increasing fluorescence. Conversely, when temperature is increased, annealed oligomers will denature/ melt , seen as decreasing fluorescence.
  To determine the rate of annealing/ melting at any given temperature, two approaches have been tried:
    The complementary oligomers are cooled from 90oC to the hold temperature being tested, say, 70oC. The ramp down rate and hold time are known (and user programmed). Once at the desired temperature, the sample is incubated for a prolonged period of time so that the change in fluorescence over time can be monitored in real time.
    The complementary oligomers are heated from 60oC to the temperature being tested, say 70oC. The ramp-up rate and the hold time are known (and user programmed). Once at the desired temperature, the sample is incubated for a prolonged period of time so that the change in fluorescence over time can be monitored in real time.
  The data in the following slides used the first protocol ie Prcl 1.

## Slide 2: Oligomer details
Name | Size | Sequence | %GC | Melting Temperature, Tm* (width/ range of melting)
A1a16 | 40bases | 5’- GACTGCAAAGATGGAAACGACCTTCTATGACGATGCCCTC-3' | 50 | 78oC (3oC)
A2 (complementary oligo) | 40bases | 5'-GAGGGCATCGTCATAGAAGGTCGTTTCCATCTTTGCAGTC-3' | 50 | 
*: This is the melting temperature in the presence of EvaGreen |  |  |  | 
Melt Curve of oligo pair A1a16-A2 performed as part of this experiment (also in the absence of dye)
Pl Note: This melt run used the same instrument parameters as the other expts. however, signal is saturated at ~ 60oC.
First derivative of the melt curve of A1a16-A2, the positive peak represents the melting temperature (Tm = 78oC, see table above).
Second derivative of the melt curve of A1a16-A2, the positive and negative peaks represent the limits of the range of temperatures over which melting occurs (width = 3oC, see table above); the point at which the plot cuts the X-axis is the Tm (78oC).
Chart: temperature (oC)

## Slide 3: Methods: Protocol 1 The experiment is carried out in the Cary Eclipse Fluorescence Spectrophotometer
Reaction Sample prepared as follows:
Reaction sample denatured as follows (program used is denat mthd 2):
  Heated up to 95oC at 10oC/min
  Held at 95oC for 1.5min
  Returned to 92oC
Reaction sample subjected to melt - kinetic run at desired temperature (74/ 75/ 76oC) (program used is prcl1 mthd 1).
Protocol 1 Mthd 1 details:
  Start thermal run at high end (90oC)
  Ramp down to desired temperature (74/ 75/ 76oC) for thermal (melt) profile
  Hold at desired temperature for kinetic run until equilibration
  See following slide for instrument parameters
Component | Source | Concentration
Native Taq Polymerase buffer | Life Technologies | 1X
MgCl2 | Life Technologies | 2mM
Oligo 1 (A1a16) | Sigma Custom Oligos | 0.5uM
Oligo 2 (A2) | Sigma Custome Oligos | 0.5uM
EvaGreen | Biotium Inc | 1X
Total vol |  | 1500ul

## Slide 4: Comparison of melting kinetics of A1a16-A2 oligo pair (40bp, %GC = 50) at 74-75-76-78oC as measured by Fluorescence Spectroscopy
Thermal (melt) run from 90 to 74/ 75/ 76/ 78/ 80oC
Kinetic run at 74/ 75/ 76/ 78/ 80oC for 600secs*
*See following slide for scale up of kinetic plots
Thermal Run Instrument Parameters |  |  | 
Start Temp |  | 90oC | 
End Temp |  | 74/ 75/ 76/ 78/ 78oC | 
Ramp Rate |  | 0.5oC/min | 
Avg Time |  | 2sec | 
Data Interval |  | 1oC | 
Kinetic Run Instrument Parameters | 
Stop Time | 10min
Signal Averaging Time (SAT) | 5secs
Data Interval | 2sec
Chart: temp (oC)
| 89.75 | 89.040000919999983 | 88.040000919999983 | 87.040000919999983 | 86.040000919999983 | 85.010002139999969 | 84.02999878 | 83.02999878 | 82.02999878 | 81.010002139999969 | 80.02999878 | 79.02999878 | 78.02999878 | 77.02999878 | 76.02999878 | 75.019996640000016 | 74.050003050000001
Int 90-74deg | 7.2564 | 7.4127 | 7.7886 | 8.099 | 8.4551 | 8.9245 | 9.3844 | 10.0438 | 11.401 | 15.1554 | 29.5632 | 77.0017 | 175.8543 | 290.4824 | 382.9025 | 453.4396 | 506.8771
Int 90-75deg | 7.6088 | 7.7571 | 8.1637 | 8.5037 | 8.9512 | 9.4605 | 10.2004 | 11.4388 | 15.6293 | 31.2251 | 82.1202 | 171.728 | 305.4834 | 401.8111 | 470.7151
Int 90-76deg | 7.3348 | 7.5754 | 7.948 | 8.347 | 8.7365 | 9.1611 | 9.7316 | 10.339 | 11.8362 | 16.4238 | 33.7605 | 88.6575 | 194.9495 | 312.6713 | 407.1006
Int 90-78deg | 5.8625 | 6.1057 | 6.3857 | 6.7483 | 7.0329 | 7.4293 | 7.8388 | 8.4511 | 9.4316 | 12.0087 | 22.3415 | 57.9256 | 141.9104 | 144.4798
Int 90-80deg | 7.3747 | 7.5927 | 7.9744 | 8.2892 | 8.771 | 9.1396 | 9.6797 | 10.3013 | 11.5085 | 14.9792 | 28.0071 | 28.7847
Chart: time (secs)
Instrument Parameters | 
Excitation wavelength | 500nm
Emission wavelength | 525nm
Excitation slit | 5nm
Emission slit | 5nm
Excitation filter | Auto
Emission filter | Open
PMT Detector Voltage | Medium (600)

## Slide 5: Summary of results
Comparison with UV Absorbance Spectroscopy
  Intensity measured is a direct measure of the amount of ds DNA, since dye is specific to ds DNA.
  Sensitivity is greater.
  Equilibration in the kinetic run takes at least 5 mins (much slower than UV spectroscopy where equilibration was seen in less than 2 mins)
  A complete thermal run (90o to 25oC) was performed using the same instrument parameters to determine melt temperature and also for purposes of calculation of the ds DNA concentration at a specific temperature. However these instrument parameters resulted in signal saturation after 60oC. At this point therefore, this calculation is not possible.
  An alternative would be to re-do the kinetics with altered instrument parameters ensuring that the sensitivity of the kinetics curves are not compromised.
The temperatures chosen for this experiment were based on the Tm of the ds oligo complex in the presence of dye (~78oC).
Expt | Thermal run |  |  | Kinetic run |  | 
 | Intensity [start] (90oC) | Intensity [end] (74/ 75/ 76/ 78/ 80oC) | Del Intensity | Intensity [start] t=0 | Intensity [end] t=120 | del Intensity
74oC | 7.2564 | 506.8771 | 499.6207 | 504.1714 | 615.1804 | 111.0090
75oC | 7.6088 | 470.7151 | 463.1063 | 470.1724 | 600.7414 | 130.5690
76oC | 7.3348 | 407.1006 | 399.7658 | 408.4270 | 562.0872 | 153.6602
78oC | 5.8625 | 144.4798 | 138.6173 | 146.5306 | 389.2415 | 242.7109
80oC | 7.3747 | 28.7847 | 21.4100 | 29.2909 | 246.3062 | 217.0153