## Slide 1: Comparative FDL fluorescence standard curve determination, with and without 5% DMSO, using Fluoroskan Ascent FL
Fluorescence standard curves were measured using the FDL Deacetylated Standard, in order to determine the appropriate concentration range of pep-Ac substrate required, for assays using the Fluoroskan Ascent FL instrument.
A concentration range of 0 – 30 uM deacetylated standard was used for standard curves without DMSO, and with 5 % DMSO.
The measurements were done using three different filter pairs: a) Ex = 355 nm/Em = 460 nm; b) Ex = 320 nm/Em = 460 nm; c) Ex = 390 nm/Em = 485 nm; in order to monitor the variations in each set (The Ex = 355 nm/Em = 460 nm pair is the closest to the recommended settings of Ex 360 nm/Em 460 nm for FDL assays).
The fluorescence signal saturated above 15 uM deacetylated standard in the assay, in case of the 355nm/460nm filter pair, thus limiting the maximum concentration of pep-Ac to 15 uM in the FDL assays, for optimum signal to noise.
The standard curves without DMSO and with 5% DMSO were found to be very similar in case of all filter pairs tested.

## Slide 2: FDL fluorescence standard curve: Ex = 355 nm / Em = 460 nm
Chart: FL Std Curve (Ex 355 nm/Em 460 nm)
| 0 | 1 | 2.5 | 5 | 10 | 15
 | 16.52 | 302.1 | 863.6 | 1876 | 2800 | 4866
 | 15.99 | 318.5 | 962.8 | 1865 | 3068 | 4766

## Slide 3: FDL fluorescence standard curve: Ex = 320 nm / Em = 460 nm
Chart: FL Std Curve (Ex 320 nm/Em 460 nm)
| 0 | 1 | 2.5 | 5 | 10 | 15
 | 16.04 | 232.6 | 664 | 1443 | 2180 | 3728
 | 15.38 | 235.1 | 711.3 | 1379 | 2273 | 3549

## Slide 4: FDL fluorescence standard curve: Ex = 390 nm / Em = 485 nm
Chart: FL Std Curve (Ex 390 nm/Em 485 nm)
| 0 | 1 | 2.5 | 5 | 10 | 15
 | 1.987 | 4.817 | 10.62 | 21.47 | 31.85 | 56.14
 | 1.939 | 5.168 | 12.3 | 22.37 | 36.86 | 58.51