## Slide 1: Validation of SIRTainty DHP1c/SIRT1 result by HPLC

## Slide 2
Experimental conditions
[SIRT1] = 5U
[NAD+] =200 mM
[H3K9 peptide] =25 mM
%DMSO = 0, 2%
[DHP1c] = 50 uM
Time points= 0, 30 min
Temp.= 37oC
HPLC-based assay
 | SIRTainty buffer |  | 2% DMSO |  | 50uM DHP1c |  | Avg | Stdv | cv%
 | no SIRT1 | yes SIRT1 | no SIRT1 | yes SIRT1 | no SIRT1 | yes SIRT1 |  |  | 
Rt_H3K9 peptide, min | 10.900 | 10.900 | 10.867 | 10.917 | 10.850 | 10.867 | 10.884 | 0.0258 | 0.24
Rt_product, min |  | 10.050 |  | 10.050 |  | 10.000 | 10.033 | 0.0289 | 0.29
DRt, min |  | 0.850 |  | 0.867 |  | 0.867 | 0.861 | 0.0098 | 1.14
Peak Are_H3K9 peptide |  | 187957 |  | 151758 |  | 149271 |  |  | 
Peak Are_product |  | 34730 |  | 31371 |  | 31281 |  |  | 
Total Area |  | 222687 |  | 183129 |  | 180552 |  |  | 
% Product |  | 15.60 |  | 17.13 |  | 17.33 |  |  | 
pmoles |  | 156.0 |  | 171.3 |  | 173.3 |  |  | 
% Activity |  |  |  | 100.0 |  | 101.1 |  |  | 
Chart: Product formed, pmoles
| Assay Buffer | 2% DMSO | 50uM DHP1c
 | 155.9588 | 171.3055 | 173.252
Remarks
HPLC results shown that the SIRT1 activity has 1.1% increase in the presence of 50uM DHP1c.
Under same experimental condition, about 80% SIRT1 activation (slides 9 and 10) by 50 uM DHP1c was detected by SIRTainty assay. However, it was not detected using HPLC.
SIRTainty assay provides false positive for those small compounds (like DHPs) who have intrinsic fluoresecence.

## Slide 3
Experimental conditions
[NAD+] =200 mM
[H3K9 peptide] =25 mM
Time points= 0, 30 min
Temp.= 37oC
__________ No 5U SIRT1
__________ Yes 5U SIRT1
Samples were collected from 96-well after read on TeCan and run on HPLC.
Except NAD+, NAM, acetyl/deacetyl H3 peptide peaks, there were other peaks from nicotinamidase, developer, which made the chromatogram very “massy”.
It was hard to identify all the peaks since the compositions of SIRTainty buffer, developer solution were unknown.
It’s difficult to identify product peaks directly from TeCan sample.
Step 1. Run samples were collected from TeCan

## Slide 4
Step 2. Run individual component included in SIRTainty assay.
SIRTainty Buffer
H3K9 peptide
NAM+
NAD+

## Slide 5
Step 3-1. Run reactions under same condition as it on TeCan- Assay buffer
Experimental conditions
[NAD+] =200 mM
[H3K9 peptide] =25 mM
Time points= 0, 30 min
Temp.= 37oC
__________ No 5U SIRT1
__________ Yes 5U SIRT1
Product
peak

## Slide 6
Step 3-2. Run reactions under same condition as it on TeCan- 2% DMSO
Experimental conditions
[NAD+] =200 mM
[H3K9 peptide] =25 mM
%DMSO= 0, 2%
Time points= 0, 30 min
Temp.= 37oC
__________ No 5U SIRT1
__________ Yes 5U SIRT1
Product
peak

## Slide 7
SIRTainty Assay
__________ 50uM DHP1c (no SIRT1)
__________ 50uM DHP1c (yes SIRT1)
Step 3-3. Run reactions under same condition as it on TeCan- 50uM DHP1c
Experimental conditions
[NAD+] =200 mM
[H3K9 peptide] =25 mM
%DMSO= 0, 2%
[DHP1c]= 0, 50uM
Time points= 0, 30 min
Temp.= 37oC
__________ No 5U SIRT1
__________ Yes 5U SIRT1
Product
peak

## Slide 8
SIRTainty assay on TeCan

## Slide 9
Repeat JMC 2015 experiment in PMC-AT Lab
Experimental conditions
[SIRT1] =5U/reaction
[NAD+] =200 mM
[H3K9 peptide] =25 mM
[DHP1c] = 50uM
% DMSO= 2%
Time points= 0, 30 min
Temp.= 37oC
Gain (TeCan )=65
 | Run 1 |  | Run 2 |  | Avg_0min | Avg_60min | D AFU | % Activity
 | 0 min | 60 min | 0 min | 60 min |  |  |  | 
Buffer only (no SIRT1) | 247 | 166 | 161 | 153 | 204.0 | 159.5 |  | 
Negative control (NAD+ peptide substrate+PNCA1+Assay buffer-no SIRT1) | 342 | 305 | 280 | 303 | 311.0 | 304.0 |  | 
Positive control (NAD+ peptide substrate+PNCA1+Assay buffer+ SIRT1) | 277 | 417 | 297 | 408 | 287.0 | 412.5 | 125.5 | 
2% DMSO (NAD+ peptide substrate+PNCA1+Assay buffer + DMSO-no SIRT1) | 315 | 233 | 275 | 253 | 295.0 | 243.0 |  | 
2% DMSO (NAD+ peptide substrate+PNCA1+Assay buffer+DMSO+SIRT1) | 276 | 347 | 250 | 312 | 263.0 | 329.5 | 66.5 | 100.0
DHP1c (NAD+ peptide substrate+PNCA1+Assay buffer+DHP1c-no SIRT1) | 1149 | 1105 | 1221 | 1099 | 1185.0 | 1102.0 |  | 
DHP1c (NAD+ peptide substrate+PNCA1+Assay buffer+DHP1c + SIRT1) | 1153 | 1234 | 1042 | 1212 | 1097.5 | 1223.0 | 125.5 | 188.7
Chart: % Activity
| 2% DMSO | 50uM DHP1c
 | 100 | 188.7
Operation procedure
The deacetylation reaction was initiated by addition of SIRT1 enzyme. The mix was incubated at 37oC for 30min. Developer was added upon the desired time point and incubated at room temperature for 30min avoiding light. Then read the sample on TeCan at Ex/Em=420nm/460nm.
Calculation details (Method A) : DAFU = AFU 60min – AFU 0min

## Slide 10
Repeat JMC 2015 experiment in PMC-AT Lab
Experimental conditions
[SIRT1] =5U/reaction
[NAD+] =200 mM
[H3K9 peptide] =25 mM
[DHP1c] = 50uM
% DMSO= 2%
Time points= 0, 30 min
Temp.= 37oC
Gain (TeCan )=65
 | Run 1 |  | Run 2 |  | Avg_0min | Avg_60min | Background Subtraction |  | D AFU | % Activity
 | 0 min | 60 min | 0 min | 60 min |  |  | DAFU_0min | DAFU_60min |  | 
Buffer only (no SIRT1) | 247 | 166 | 161 | 153 | 204.0 | 159.5 |  |  |  | 
Negative control (NAD+ peptide substrate+PNCA1+Assay buffer-no SIRT1) | 342 | 305 | 280 | 303 | 311.0 | 304.0 |  |  |  | 
Positive control (NAD+ peptide substrate+PNCA1+Assay buffer+ SIRT1) | 277 | 417 | 297 | 408 | 287.0 | 412.5 | -24.0 | 108.5 | 132.5 | 
2% DMSO (NAD+ peptide substrate+PNCA1+Assay buffer + DMSO-no SIRT1) | 315 | 233 | 275 | 253 | 295.0 | 243.0 |  |  |  | 
2% DMSO (NAD+ peptide substrate+PNCA1+Assay buffer+DMSO+SIRT1) | 276 | 347 | 250 | 312 | 263.0 | 329.5 | -32.0 | 86.5 | 118.5 | 100.0
DHP1c (NAD+ peptide substrate+PNCA1+Assay buffer+DHP1c-no SIRT1) | 1149 | 1105 | 1221 | 1099 | 1185.0 | 1102.0 |  |  |  | 
DHP1c (NAD+ peptide substrate+PNCA1+Assay buffer+DHP1c + SIRT1) | 1153 | 1234 | 1042 | 1212 | 1097.5 | 1223.0 | -87.5 | 121.0 | 208.5 | 175.9
Chart: % Activity
| 2% DMSO | 50uM DHP1c
 | 100 | 175.9
Calculation details (Method B) :
  DAFU = DAFU_60min – DAFU_0min
  DAFU_0min = AFU(with SIRT1) – AFU(without SIRT1) at 0min time point
  DAFU_60min =AFU(with SIRT1) – AFU(without SIRT1) at 60min time point
Operation procedure
The deacetylation reaction was initiated by addition of SIRT1 enzyme. The mix was incubated at 37oC for 30min. Developer was added upon the desired time point and incubated at room temperature for 30min avoiding light. Then read the sample on TeCan at Ex/Em=420nm/460nm.